SELF-ASSESSMENT - MODULE 3-5: Gas Movement

- INFORMATION NEEDED TO CALCULATE PATIENT FLOW NEEDS

  - CALCULATIONS:

    - Weight in kg=Weight in pounds2.2

    - VE=Vt×f

    - Vt=VEf

    - f=VEVt

    - Patient peak inspiratory flow demand (PIF) = VE×(I+E)

    - Patient peak inspiratory flow demand (PIF) = VttI×60

  - ADULT NORMAL VALUES:

    - Inspiratory Time: 0.8 – 1.2 seconds

    - Tidal Volume: spontaneous 5 – 8 mL/kg of IBW

    - Respiratory Rate: 10 – 20 breaths/min

    - Minute volume: 5 – 10 L/min

    - I : E Ratio: 1 : 2 to 1 : 4

    - Normal range for adult inspiratory flow: 24 – 30 L/min, but may be as high as 60-100 L/min.

  - EXERCISES:

    - A patient’s weight is measured as 150 pounds.   What is the range of normal tidal volumes?

      - Weight in kg=Weight in pounds2.2=150lbs2.2lbskg=68.18kg≈68.2kgTidal Volume=5-8 mLkg68.2 kg×5mLkg=341 mL   68.2 kg×8mLkg=545 mL

    - A patient’s weight is measured as 200 pounds.   What is the range of normal tidal volumes?

      - Weight in kg=Weight in pounds2.2=200lbs2.2lbskg=90.9kgTidal Volume=5-8 mLkg90.9 kg×5mLkg=455 mL   90.9 kg×8mLkg=727 mL

    - A patient’s weight is measured as 40 kilograms.   What is the range of normal tidal volumes?

      - Tidal Volume=5-8 mLkg40 kg×5mLkg=200 mL  40 kg×8mLkg=320 mL

    - A patient has a tidal volume of 600 mL and a frequency of 20/min.  Calculate the minute ventilation (VE).

      - VE=Vt×f=600mLbreath×20breathsmin=12,000mLmin=12Lmin

    - A patient has a tidal volume of 350 mL and a frequency of 15/min.  Calculate the minute ventilation (VE).

      - VE=Vt×f=350mLbreath×15breathsmin=5,250mLmin=5.3Lmin

    - A patient has a VE of 12.5 L/min and a frequency of 25/min.  Calculate the average tidal volume.

      - Vt=VEf=12.5Lmin25breathsmin=0.5Lbreath

    - A patient has a VE of 8.4 L/min and a frequency of 14/min.  Calculate the average tidal volume.

      - Vt=VEf=8.4Lmin14breathsmin=0.6Lbreath

    - A patient has a VE of 10 L/min and a tidal volume of 500 mL.  Calculate the frequency.

      - f=VEVt=10Lmin500mLbreath=10Lmin0.5Lbreath=20breathsmin

    - A patient has a VE of 5.8 L/min and a tidal volume of 400 mL.  Calculate the frequency.

      - f=VEVt=10Lmin500mLbreath=10Lmin0.5Lbreath=20breathsmin

    - If the patient’s Vt is 550 mL, and the inspiratory time is 0.9 seconds, calculate the patients peak inspiratory flow.

VttI=550mL0.9sec×60=0.55L0.9sec×60secmin=.611×60=36.7Lmin

      - PIF =

    - If the patient’s Vt is 680 mL, and the inspiratory time is 1.2 seconds, calculate the patients peak inspiratory flow.

VttI=680mL1.2sec×60=0.68L1.2sec×60secmin=.567×60=34Lmin

      - PIF =

    - If the patient’s VE is 10.5 L/min, the f is 12/min, and the inspiratory time is 1.0 seconds, calculate the peak inspiratory flow.

Vt=VEf=10.5Lmin12breathsmin=0.875LbreathVttI=0.88L1.2sec×60secmin=.733×60=44Lmin

      - PIF  =

    - If the VE is 7.6 L/min, the f is 16/min, and the inspiratory time is 0.8 seconds, calculate the peak inspiratory flow

Vt=VEf=7.6Lmin16breathsmin=0.475LbreathVttI=0.48L0.8sec×60secmin=.6×60=36Lmin

      - PIF =

    - VE=Vt×f=500mLbreath×20breathsmin=10,000mLmin=10.0LminGiven a frequency of 20/min and a tidal volume of 500 mL, calculate the patient’s minute ventilation.

    - VE=Vt×f=600mLbreath×15breathsmin=9,000mLmin=9.0LminGiven a frequency of 15/min and a tidal volume of 600 mL, calculate the patient’s minute ventilation.

    - VE=Vt×f=800mLbreath×25breathsmin=20,000mLmin=20.0LminGiven a frequency of 25/min and a tidal volume of 800 mL, calculate the patient’s minute Ventilation.

    - Given a minute ventilation of 7.5 L/min and a frequency of 16/min, calculate the average tidal volume.

      - Vt=VEf=7.5Lmin16breathsmin=0.469Lbreath=469mLbreath

    - Given a minute ventilation of 10. 0 L/min and a frequency of 12/min, calculate the average tidal volume.

      - Vt=VEf=10Lmin12breathsmin=0.833Lbreath=833mLbreath

    - Given a minute ventilation of 12.5 L/min and a frequency of 21, calculate the average tidal volume.

      - Vt=VEf=12.5Lmin21breathsmin=0.595Lbreath=595mLbreath

    - Given a minute ventilation of 8.4 L/min and a tidal volume of 600 mL, calculate the frequency.

      - f=VEVt=8.4Lmin600mLbreath=8.4Lmin0.6Lbreath=14breathsmin

    - Given a minute ventilation of 8.8 L/min and a tidal volume of 550 mL, calculate the frequency.

      - f=VEVt=8.8Lmin550mLbreath=8.8Lmin0.55Lbreath=16breathsmin

    - Given a minute ventilation of 6.4 L/min and a tidal volume of 400 mL, calculate the frequency.

      - f=VEVt=6.4Lmin400mLbreath=6,4Lmin0.4Lbreath=16breathsmin

    - Given an inspiratory time of 1.2 seconds and a peak inspiratory flow of 35 L/min, calculate the tidal volume.

      - PIF=VttI×60,PIF×tI60=Vt=35Lmin×1.2sec60minsec=0.7L=700mL

    - Given an inspiratory time of 1.0 second and a peak inspiratory flow of 28 L/min, calculate the tidal volume.

      - PIF=VttI×60,PIF×tI60=Vt=28Lmin×1.0sec60minsec=0.467L=467mL

    - Given a tidal volume of 450 mL and an inspiratory time of 1.4 seconds, calculate the peak inspiratory flow.

      - PIF=VttI×60=450mL1.4 sec×60secmin=19,286  mLmin=19.3Lmin

    - Given a tidal volume of 625 mL and an inspiratory time of 0.8 seconds, calculate the peak inspiratory flow.

      - PIF=VttI×60=625 mL0.8 sec×60secmin=46,875mLmin=46.9Lmin

    - Given a minute ventilation of 12 L/min, and a I:E ratio of 1:3, calculate the minimal inspiratory flow needed to meet the patients inspiratory needs.

      - PIF = VE×(I+E)=12Lmin×1+3=12Lmin×4=48Lmin

    - Given a minute ventilation of 8.6 L/min, and a I:E ratio of 1:2, calculate the minimal inspiratory flow needed to meet the patients inspiratory needs.

      - PIF = VE×(I+E)=8.6Lmin×1+2=8.6Lmin×3=25.8Lmin

- CALCULATING SYSTEM TOTAL FLOW

  - An air-entrainment nebulizer is set at an FIO2 of 0.40 and the oxygen flowmeter is set at 8 liters/min.  Calculate the following:

air:oxygen=100−FIO2FIO2−21=100−4040−21=6019=3.2≈3:1

    - Air:O2 ratio:

    - O2 liter Flow: 8 L/min

AirFlow=Air Ratio×OxygenFlow=3×8Lmin=24Lmin

    - Air Liter Flow:

    - Total Liter Flow: Oxygen Flow + Air Flow = 8 L/min + 24 L/min = 32 L/min

  - The air-entrainment mask is set at an FIO2 of 0,28 and the oxygen flowmeter is set at 3 liters/min.  Calculate the following:

air:oxygen=100−FIO2FIO2−21=100−2828−21=727=10.3≈10:1

    - Air:O2 ratio:

    - AirFlow=AirRatio×OxygenFlow=10×3Lmin=30LminO2 liter Flow: 3 L/min

    - Air Liter Flow:

    - Total Liter Flow: Oxygen Flow + Air Flow = 3 L/min + 30 L/min = 33 L/min

  - An air-entrainment nebulizer is set at an FIO2 of 0.70 and the oxygen flowmeter is set at 6 liters/min.   Calculate the following:

air:oxygen=100−FIO2FIO2−21=100−7070−21=3049=0.61≈0.6:1

    - Air:O2 ratio:

    - O2 liter Flow: 6 L/min

AirFlow=AirRatio×OxygenFlow=0.6×6Lmin=3.6Lmin

    - Air Liter Flow:

    - Total Liter Flow: Oxygen Flow + Air Flow = 6 L/min + 3.6 L/min = 9.6 L/min

  - An air-entrainment mask is set at an FIO2 of 0.50 and the oxygen flowmeter is set at 8 liters/min.  Calculate the following:

air:oxygen=100−FIO2FIO2−21=100−5050−21=5029=1.7:1

    - Air:O2 ratio:

    - O2 liter Flow: 8 L/min

AirFlow=AirRatio×OxygenFlow=1.7×8Lmin=13.6Lmin

    - Air Liter Flow:

    - Total Liter Flow: Oxygen Flow + Air Flow = 8 L/min + 13.6 L/min = 21.6 L/min

  - An air-entrainment nebulizer is set at an FIO2 of 1.0 and the oxygen flowmeter is set at 15 liters/min.  Calculate the following:

air:oxygen=100−FIO2FIO2−21=100−100100−21=079=0:1

    - Air:O2 ratio:

    - O2 liter Flow: 15 L/min

AirFlow=AirRatio×OxygenFlow=0×15Lmin=0Lmin

    - Air Liter Flow:

    - Total Liter Flow: Oxygen Flow + Air Flow = 15 L/min + 0 L/min = 15 L/min

  - Assuming the flowrate stays the same on an air-entrainment device, what happens to total liter flow as the FIO2 increases? IT GOES DOWN

  - The concentration of oxygen delivered by an air-entrainment system can be varied by:

    - Altering the size of the jet orifice

    - Altering the size of the air entrainment ports

    - Both 1 and 2.

  - Backpressure on an air-entrainment system decreases the volume of fluid or gas entrained.  This causes the oxygen concentration delivered by the system to ____________.

    - Increase

    - Decrease

    - Stay the same

# PATIENT NEEDS AND DEVICE DELIVERY

  - You are setting up an air-entrainment mask at an FIO2 of 0.40 and the oxygen flowmeter is set at 12 l/min.  The patient’s tidal volume is 600 mL and the inspiratory time is 1.5 seconds.  Is the flow from this system meeting the patient’s inspiratory needs?

air:oxygen=100−FIO2FIO2−21=100−4040−21=6019=3.2≈3:1

    - Air: Oxygen Ratio:

    - Total Liter Flow: Oxygen Flow + Air Flow = 12 L/min + 36 L/min = 48 L/min

VttI×60=0.6L1.5 sec×60sec min=0.4×60=24Lmin

    - Peak Inspiratory Flowrate:

    - Is the FDO2 > FIO2? YES  NO

    - What FIO2 would the patient actually receive?

      - 0.40

      - Less than 0.40

      - Greater than 0.40

  - You are setting up an air-entrainment nebulizer with an aerosol mask at an FIO2 of 0.70 and the oxygen flowmeter is set at 12 L/min.  The patient’s minute ventilation is 8 L/min, the inspiratory time is 0.5 seconds, and the respiratory rate is 10/min.  Is the flow from this system meeting the patient’s inspiratory needs?

air:oxygen=100−FIO2FIO2−21=100−7070−21=3049=0.61=0.6:1

    - Air: Oxygen Ratio:

    - Vt=VEf=8Lmin10breathsmin=0,8LbreathVttI×60=0.8L0.5 sec×60sec min=1.6×60=96Lmin Total Liter Flow: Oxygen Flow + Air Flow = 12 L/min + 7.2 L/min = 19.2 L/min

    - Peak Inspiratory Flowrate:

    - Is the FDO2 > FIO2?  YES   NO

    - What FIO2 would the patient actually receive?

      - 0.70

      - Less than 0.70

      - Greater than 0.70

  - You are setting up an air-entrainment nebulizer with a tracheostomy mask at an FIO2 air:oxygen=100−FIO2FIO2−21=100−3535−21=6514=4.6≈5:1of 0.35 and the oxygen flowmeter is set at 15 L/min.  The patient’s minute ventilation is 8 L/min and the I:E ratio is 1:3.  Is the flow from this system meeting the patient’s inspiratory needs?

    - Air: Oxygen Ratio:

    - VE×(I+E)=8Lmin×1+3=8Lmin×4=32LminTotal Liter Flow: Oxygen Flow + Air Flow = 15 L/min + 75 L/min = 90 L/min

    - Peak Inspiratory Flowrate:

    - Is the FDO2 > FIO2?  YES   NO

    - What FIO2 is the patient actually receiving?

      - 0.35

      - Less than 0.35

      - Greater than 0.35

  - air:oxygen=100−FIO2FIO2−21=100−6060−21=4039=1:1You are setting up an air-entrainment nebulizer with a Briggs (t-) adapter at an FIO2 of 0.60 and the oxygen flowmeter is set at 12 L/min.  The patient’s tidal volume is 400 mL and the inspiratory time is 0.9 seconds.  Is the flow from this system meeting the patient’s inspiratory needs?

    - Air: Oxygen Ratio:

    - VttI×60=0.4L0.9 sec×60sec min=0.44×60=26.7Lmin Total Liter Flow: Oxygen Flow + Air Flow = 12 L/min + 12 L/min = 24 L/min

    - Peak Inspiratory Flowrate:

    - Is the FDO2 > FIO2?  YES   NO

    - What FIO2 would the patient receive?

      - 0.60

      - Less than 0.60

      - Greater than 0.60

  - VE×(I+E)=6.8Lmin×1+1.5=6.8Lmin×2.5=17.0LminGiven a minute ventilation of 6.8 L/min and a I:E ratio of 1:1.5, calculate the minimal inspiratory flow needed to meet the patient’s inspiratory needs.

    - PIF:

    - air:oxygen=100−FIO2FIO2−21=100−4040−21=6019=3.2≈3:1The doctor has ordered an air-entrainment mask set at an FIO2 of 0.40 and the oxygen flowmeter is set at 6 L/min.  Is the total flowrate from this system sufficient to meet the patient’s inspiratory needs?

Oxygen Flow + Air Flow = 6 L/min + 18 L/min = 24 L/min

    - What will happen to the FIO2 we are giving the patient?  THE DEVICE’S TOTAL FLOW EXCEEDS THE PATIENT’S INSPIRATORY FLOW RATE SO THE DESIRED FIO2 WILL BE DELIVERED.

  - VE×(I+E)=11Lmin×1+2=11Lmin×3=33LminGiven a minute ventilation of 11 L/min and an I:E ratio of 1:2, calculate the minimal inspiratory flow needed to meet the patients inspiratory needs.

    - PIF:

    - air:oxygen=100−FIO2FIO2−21=100−6060−21=4039=1:1The doctor has ordered an air-entrainment nebulizer with an aerosol mask at an FIO2 of 0.60 and the oxygen flowrate is set at 10 L/min.  Is the total flowrate from this system sufficient to meet the patient’s inspiratory needs?

Oxygen Flow + Air Flow = 6 L/min + 18 L/min = 24 L/min

    - What will happen to the FIO2 we are giving the patient?  THE DEVICE’S TOTAL FLOW DOES NOT EXCEED THE PATIENT’S INSPIRATORY FLOW RATE SO THE DESIRED FIO2 WILL NOT BE DELIVERED.

- OXYGEN DELIVERY SYSTEMS

  - According to Egan, there are four categories of oxygen delivery systems,

    - LOW-FLOW DEVICES

    - HIGH-FLOW DEVICES

    - RESERVOIR SYSTEMS

    - ENCLOSURES

  - The most common low-flow system is the NASAL CANNULA.

    - This system should be run between 1/4 & 8 liters per minute oxygen flow and will deliver approximately 24 - 40 % oxygen.

    - The oxygen concentration of this system depends on the patients respiratory pattern.  As the patient begins to breathe more deeply and rapidly, the oxygen concentration will go (up or down).

  - Even though reservoir systems run at close to the same flow rates of oxygen, the concentrations of oxygen provided are higher.  Why is this?  RESERVOIR SYSTEMS INCORPORATE A MECHANISM FOR GATHERING AND STORING OXYGEN BETWEEN PATIENT BREATHS.

  - Why is the oxygen concentration with most low-flow, reservoir and enclosure systems variable?  THE TOTAL FLOW IS LESS THAN THE PATIENT’S INSPIRATORY FLOW RATE.

  - What do all high flow systems have in common?  TOTAL FLOW EXCEEDS THE PATIENT’S INSPIRATORY FLOW RATE.

  - Name four different systems or set ups that will provide a fixed oxygen concentration.

    - AIR-ENTRAINMENT MASKS

    - AIR-ENTRAINMENT NEBULIZERS

    - BLENDERS

    - DUAL FLOWMETERS

  - Scenario:  You are called to set up oxygen on a patient in the Emergency Department.  You are told the patient is 72 years old with a history of emphysema.  He is in obvious respiratory distress (respiratory rate 28, accessory muscle use and bilateral wheezing) and his oximetry (SpO2) on room air is 86%.  You decide you would like to begin at approximately 30% oxygen.  Which oxygen delivery system(s) would be appropriate?

    - AIR-ENTRAINMENT MASK TO DELIVER A PRECISE CONCENTRATION IN THE FACE OF A VARIABLE RESPIRATORY RATE AND PATTERN.

  - What is the formula for calculating minute ventilation (VE)?  VE=Vt×f

  - What is the formula for calculating peak inspiratory flow (PIF)? 
 
Patient peak inspiratory flow demand (PIF) = VE×(I+E)or VttI×60

  - What is the calculation for normal spontaneous tidal volume (Vt)? 5 to 8 mL/kg IBW

  - Convert 180 pounds to kilograms.   Weightinkg=Weight in pounds2.2Weightinkg=180lbs2.2=81.8kg

  - What is the frequency of a person with a minute ventilation of 10 L/min and a tidal volume of 500 mL?  f=VEVt=10Lmin500mLbreath=10Lmin0.5Lbreath=20breathsmin

  - What is the tidal volume of someone with a minute ventilation of 8 L/min and a frequency of 15 breaths/minute?  
Vt=VEf=8Lmin15breathsmin=0.533Lbreath=533mLbreath

  - air:oxygen=100−FIO2FIO2−21=100−7575−21=2554=0.46≈0.5:1You have a nebulizer and aerosol mask set up at an FIO2 of 0.75.  Calculate the air:oxygen entrainment ratio.

    - AirFlow=AirRatio×OxygenFlow=0.5×10Lmin=5LminIf you run your oxygen flow at 10 L/min, what will the air entrainment be?

    - What is the total flow provided to the aerosol mask?

Total Flow = Oxygen Flow + Air Flow = 10 L/min + 5 L/min = 15 L/min

    - Your patient needs 30 L/min flow.  Are you delivering enough flow with your device to meet their inspiratory needs?  YES  NO

    - If not, what will happen to the inspired oxygen concentration?  IT WILL BE REDUCED.

    - If not, what can you do to correct the situation?  USE AN ALTERNATE SYSTEM LIKE A BLENDER OR A DUAL FLOWMETER SYSTEM.

  - I have a patient who needs 16 L/min inspiratory flow and an FIO2 of 0.40. I have an air and oxygen flowmeter and want to mix (blend) these gases together in the proper ratios and flows to meet my patient’s needs.

air:oxygen=100−FIO2FIO2−21=100−4040−21=6019=3.2≈3:1

    - Where should I set the oxygen flowmeter?  4 L/min

    - Where should I set the air flowmeter?  12 L/min

  - If an oxygen flowmeter is set at 10 L/min and an air flowmeter is set at 17 L/min, what is the oxygen concentration being delivered?  

FDO2=Oxygen Flow Rate+(.2×Air Flow Rate)Total FlowFDO2=10+.2×1727=10+3.427=13.427=.496≈0.50

  - Which states of matter are considered fluids?

    - GASES

    - LIQUIDS

  - Define flow and give an example of one of its unit of measure.

    - Definition: THE BULK MOVEMENT OF A SUBSTANCE THROUGH SPACE EXPRESSED AS VOLUME OF FLUID MOVED PER UNIT OF TIME.

    - Unit of measure: LITERS PER MINUTE, LITERS PER SECOND

  - Pressure is defined as force/specific surface area.  For a static fluid, pressure is dependent on VELOCITY x CROSS-SECTIONAL AREA.

  - VISCOSITY is the property of a fluid that opposes flow.

  - Three patterns of fluid flow are:

    - LAMINAR

    - TURBULENT

    - TRANSITIONAL

  - Poiseuille’s Law describes the factors effecting laminar flow.  Write the formula below and define the variables.

P=V8ℓnπr4P = Driving pressure (to move gas through a tube)–fluid viscosity (n)–tube length (l)–flow (?)–radius (r)

  - If viscosity of gas or the length of the tube increases, what happens to driving pressure?  (if flow is to remain the same)  INCREASES

  - If the radius of the tube decreases, what happens to driving pressure?  (if flow is to remain the same)  INCREASES

    - In what situation might this apply clinically?  BRONCHOSPASM

  - A fluids flow becomes turbulent when the Reynolds number is > 2,000

  - Reynold's Number=ν×d×2rhv = linear velocity (distance/time)d = fluid density (weight/volume)r  = tube radius (size of opening)h = fluid viscosity (thickness, stickiness)What is the formula for Reynolds number?  Define the variables.

  - If fluid velocity, fluid density or tube radius go up, the Reynolds number will INCREASE.

  - If fluid viscosity goes down, Reynolds number will DECREASE.

  - What is the Bernoulli Effect?  AS FLUID FLOWS THROUGH A TUBE AND MEETS A RESTRICTION, THE FORWARD VELOCITY WILL INCREASE AND THE LATERAL WALL PRESSURE WILL DECREASE.

  - Fluid velocity at a constant flow varies inversely with its LATERAL WALL PRESSURE.

  - What is the Venturi Principle?  IF GAS FLOWING THROUGH A TUBE MEETS A SMALL ENOUGH CONSTRICTION, THE PRESSURE WILL DROP TO SUB ATMOSPHERIC AND ACTUALLY ENTRAIN A SECOND GAS (FLUID).

  - Fluids have three kinds of energy

    - POTENTIAL

    - KINETIC

    - PRESSURE

  - Gravity increases the effect of POTENTIAL energy.

  - KINETIC energy is the result of fluid in motion (velocity).

  - PRESSURE energy is the lateral force exerted by moving fluid on the walls of its container.

  - One of the Laws of Thermodynamics states that the energy at any point in a fluid stream is the same through the stream where energy = Velocity or Kinetic Energy x Lateral Pressure Energy.  As a tube narrows, the velocity will INCREASE and the lateral pressure will DECREASE.

  - This Bernoulli Effect will allow for fluid ENTRAINMENT at the point of narrowing.

  - Entrainment with an air injector is dependent on the size of the

    - JET ORIFICE OPENING

    - AIR ENTRAINMENT PORT

  - Clinical examples of Respiratory Therapy equipment that uses this theory are

    - AIR-ENTRAINMENT MASKS

    - AIR-ENTRAINMENT NEBULIZERS

  - A clinical example of Respiratory Therapy equipment that uses the Bernoulli Effect is AIR-ENTRAINMENT NEBULIZERS.

  - Pressure past the narrowing point can be almost completely restored if the angle of the tube dilation does not exceed 15 degrees.

  - A smaller jet or larger entrainment port will allow GREATER air entrainment.