PERFORMANCE EVALUATION #45		STUDENT NAME:  ________________________

Automatic Targeting Scheme (Dual mode)

INSTRUCTOR NAME:  _____________________

(Update 11/2015)

GRADE: Pass Fail (Circle One)

Directions:  This exercise could be done with the NPB 840, Servo i or Drager XL ventilators.  Test lungs will be set-up by the instructor.  Patient parameters will be provided by instructor.

| Points | # PERFORMANCE EVALUATION SCORING |
| --- | --- |
| 3 | Describes and/or performs objectives perfectly without prompting and in appropriate time interval. |
| 2 | Describes and/or performs objectives satisfactorily with minimal prompting or assistance/or completes step slower than expected. |
| 1 | Describes and/or performs objectives with assistance or prompting. Appears unsure of task. |
| 0 | Unable to or fails to perform objective adequately. |
| NA | Objective not appropriate or unnecessary. |

Students must pass all critical steps (*) with a score of 2 or 3

PATIENT SCENARIOS

Start Time __________________   	End Time _____________________

Instructor will circle one:  	Airway information:   8.0 OET   or   8.0 Trach

Instructor will circle one:  	Humidification:  HME  or  Heated wire  or Non-heated wire

Instructor will circle one:

| Diagnosis | COPD | ARDS | Post-op CABG (Normal Lung) | Asthma |
| --- | --- | --- | --- | --- |
| FiO2 | FiO2 0.21 | FiO2 0.21 | FiO2 0.21 | FiO2 0.21 |
| PEEP | 0 | 0 | 0 | 0 |
| Machine Volume | 500 | 400 | 600 | 500 |
| Spon. Volume | 300 | 200 | 400 | 200 |
| Ti | 1.0 | 1.4 | 1.2 | 0.8 |
| Rate | 10 | 15 | 12 | 10 |

Instructor will set:

|  | Cs | Raw | Vt/ Vtmax% | V/ Vmax% | Rate | Offset |
| --- | --- | --- | --- | --- | --- | --- |
| Red Test Lung | 0.0.4 to 0.6 | 5/20 |  |  |  |  |
| Blue Test Lung No spon. breaths | 0 spring | 2 |  |  |  |  |
| Blue Test Lung spon. breaths | 0 spring | 2 | 30 | 15 | 15 | 4 |

Instructor will circle one:

| Ventilator | NPB 840 | Drager XL | Servo i |
| --- | --- | --- | --- |
| Automatic Targeting Scheme | AC/VC+ | CMV (Autoflow on) | AC/PRVC |
|  | VS | --- | VS |

|  | 0 | 1 | 2 | 3 | NA |
| --- | --- | --- | --- | --- | --- |
| - *Initial settings input properly. |  |  |  |  |  |
| - Alarm or other problem is detected and corrected as needed. |  |  |  |  |  |
| - Leaks are detected properly and fixed. |  |  |  |  |  |
| - *Apnea settings input properly. |  |  |  |  |  |
| - *Alarms set properly. |  |  |  |  |  |
| - *“Other settings” input properly (depending on ventilator). |  |  |  |  |  |
| Time for set up: ________________ minutes. (expected time 20 minutes or less) |  |  |  |  |  |
| Actual pressure reading: _______cm H2O. |  |  |  |  |  |
| - Can evaluate graphic display and explain how they know the slope (rise) is set up properly. |  |  |  |  |  |
| - Can explain purpose of high pressure alarm setting in this mode. |  |  |  |  |  |
| - High pressure alarm setting ________ cm H2O |  |  |  |  |  |
| - Can explain purpose of High Tidal Volume alarm in this mode. |  |  |  |  |  |
| - High Tidal Volume alarm setting ________ mL. |  |  |  |  |  |
| - Can evaluate graphic display and explain how they know the ESENS% or Insp. Cycle Off% is set up properly. |  |  |  |  |  |
| - Can demonstrate how to freeze and analyze graphics. |  |  |  |  |  |

Instructor will change:     1. Red test lung compliance		2.  Blue test lung resistance

|  | 0 | 1 | 2 | 3 | NA |
| --- | --- | --- | --- | --- | --- |
| - Can explain why alarms are sounding. |  |  |  |  |  |
| - Can explain what the ventilator is doing in response to change. |  |  |  |  |  |
| - Can describe what can be done to correct the situation. |  |  |  |  |  |
| - Answers oral review questions correctly. |  |  |  |  |  |

Instructor Comments:  __________________________________________________________

_____________________________________________________________________________

ORAL REVIEW QUESTIONS:

- What is the predicted tidal volume range when ventilating a patient with normal lungs?

5 – 8 mL/Kg IBW

- What is the predicted tidal volume range when ventilating a patient with obstructive disease such as COPD or asthma?

5 – 8 mL/Kg IBW

- What is the predicted tidal volume range when ventilating a patient with non-compliant lungs such as ARDS?

6 - 8 mL/Kg IBW and most severe may be less than 6 mL/Kg IBW (maintain Pplat 30 cm H2O or less)

- What is the desired tidal volume range when a ventilator patient is breathing spontaneously?

5 – 8 cc/Kg IBW

- What is the allowable variation between set volume on the ventilator and exhaled (returned) volume by the patient?

10% otherwise the leak must be corrected (circuit or airway) or documented if not correctable (uncuffed airway or chest tube)

- What is the appropriate setting for Apnea Back-up rate?

Set the rate and volume to support the patient at a similar minute ventilation should they stop breathing.

- If the apnea interval is set at 20 seconds, how low does the rate need to be before the apnea alarm will go off?

60 / 20 = 3 breaths per minute therefore it is rare for apnea ventilation to occur in SIMV or AC and more common in CPAP, ATC and Pressure Support.

- What type of things are set on the NPB 840 when your access the “More Settings” area?

Humidifier type, O2 analyzer and D sens%

- What is D sens% for?

It allows the practitioner to determine how much leak must occur before the disconnect alarm is activated.  Default setting is 75%.  This means a leak of 75% of the set volume must occur before the disconnect alarm will sound.

- What is the purpose of rise time?

It allows the practitioner to adjust how quickly the inspiratory pressure will be reached.

- When can rise time be applied?

Pressure ventilation (mechanical or pressure supported breaths)

- When should rise time be decreased (Drager & Vision shorter time NPB 840 higher %)?

When a scoop is observed in the beginning of the pressure-time curve.  The curve should be square at the beginning.

- When should rise time be increased (Drager & Vision longer time NPB 840 lower %)?

When a spike is observed in the beginning of the pressure-time curve.  The curve should be square at the beginning.

- What mode of ventilation “flow-cycles”?

Pressure support and Volume Support

- What is the purpose of E sens% or Inspiratory Cycle off% ?

To allow the practitioner to adjust the point where a pressure supported or volume supported breathes will flow cycle.

- When is E sens% or Inspiratory Cycle off% applied?

Pressure supported & Volume supported breaths

- When should E sens% or Inspiratory Cycle off%  be decreased?

For patients that desire a longer inspiratory time such as stiff lung patients (ARDS) and for patients where there is a scoop at the end of the breath or a double cycle.

- When should E sens% or Inspiratory Cycle off% be increased?

For patients that desire a shorter inspiratory time such as obstructive lung patients (COPD) and for patients where there is a spike at the end of the breath or air trapping.

- In Automatic Targeting Scheme modes of ventilation, what does the ventilator do in response to changes in compliance and resistance?

The ventilator will titrate the pressure up or down 1 – 3 cm H2O per breath in an attempt to deliver the set volume.

- What can be done if the volume cannot be delivered with the maximum allowable pressure?

The therapist should find the underlying cause.  They can check the breath sounds and airway, suction, bronchodilate, suggest diuresis, increase the pressure, decrease the volume…

- What will happen if a patient in Volume Support goes apneic?

Most ventilators have some sort of backup ventilation.  Be sure to know how each works.  Some return the patient to volume ventilation (Servo i).

- What happens if an apneic patient begins to breathe again?

Some ventilators will return automatically to VS (Servo i).  Be sure you know what each one does.

- On the Drager XL, what occurs when the autoflow is turned on.

Ventilator goes into Automatic Targeting Scheme ventilation (pressure ventilation with volume target)

- What is the purpose of Pmax on the Drager XL

Pmax allows you to limit the pressure in volume ventilation.