# RSPT 2350: MODULE K - CAPNOGRAPHY

- DEFINITIONS AND ABBREVIATIONS

  - ETCO2 - End Tidal CO2

  - PETCO2 - End Tidal CO2 tension

  - a-ADCO2 - The difference between arterial and alveolar (end-tidal) CO2

    - P(a-ET)CO2

    - PACO2 approximates PETCO2

  - PECO2 - Mean Expired CO2.

    - This value is used in the calculation of the Bohr Equation.

      - Vd=PaCO2−PECO2PaCO2×Vt

Vd=40−2840×500mL=150mL

      - VdVt=PaCO2−PECO2PaCO2

VdVt=40−2840=.3=30%

  - Capnometry - is the measurement and numerical display of CO2 appearing at the airway.

  - Capnography - is the measurement and graphic display of the CO2 appearing at the airway.

    - The device used to measure CO2 and displays a waveform is called a capnograph.

    - Capnography includes capnometry.

  - Capnogram - the waveform displayed by a capnograph.

  - Mass Spectrometry - is the measurement of CO2 by which CO2 is separated in a gas mixture by its mass and the number of ionized molecules is measured to determine the concentration in the gas mixture.  A mass spectrometry can be used to measure CO2, O2, N2, and anesthetic agents.

  - Infrared Spectrometry - compares the amount of infrared light absorbed by a sample of patient gas to the amount of light absorbed in a chamber containing no carbon dioxide.  Methods of sampling include:

    - Mainstream Sampling - The optical bench is attached to an adapter at the patient’s airway.

    - Sidestream Sampling - The respiratory gases are continuously aspirated from an adaptor placed at the patient’s airway and transported to the optical bench via sample tubing.

    - Most bedside capnographs used in respiratory care measure CO2 by infrared absorption.  The adapter must be placed between the ventilator wye and the ET tube.  If an HME is used, the airway adapter should be placed between the ET tube and the HME.

- CONCENTRATIONS OF CO2

  - Atmosphere  (PICO2)

    - 0.04%

    - 0.3 mm Hg  (Barometric pressure x .0004)

  - Alveoli

    - 5.6%

    - 40 mm Hg

  - Arterial Blood

    - 5.6%

    - 40 mm Hg

  - Exhaled

    - PETCO2 should approximate alveolar in the absence of lung disease.  This can be measured breath by breath with a capnograph.

    - PETCO2 is the mean expired CO2 and is approximately 28 torr

    - This must be measured by taking a collection of expired gas over a minute and calculating the mean value (Douglas Bag and a capnograph).

    - P(a-ET)CO2 gradient is normally less than 5 mm Hg.

- CO2 WAVEFORM AND TREND RECORDING

  - Axis

    - Vertical axis is CO2 in % or mm Hg.

    - Horizontal axis is time.

  - Phases of the CO2 waveform

    - Phase I  (A-B) represents deadspace ventilation (minimal CO2).

    - Phase II (B-C) represents mixed deadspace ventilation and alveolar ventilation.

      - Sharp rise in CO2 is noted.

    - Phase III (C-D) represents the gas coming from the alveoli and is called the alveolar plateau.

      - Little change in CO2 level.

    - Phase IV (D) is the PETCO2.

  - Airflow obstruction (secretions, COPD, partially kinked ET tube, bronchospasm) will result in an absence of phase III (alveolar plateau or a rise alveolar plateau).

  - Measuring Mean Expired CO2 (PECO2).

    - The mean expired CO2 is the mean value calculated from all 4 phases of the CO2 capnogram.

- PETCO2 AS A REFLECTION OF CIRCULATION, METABOLISM AND VENTILATION

  - Metabolism

    - CO2 is produced in the body as an end product of food (energy) metabolism.

    - C6H12O6 + 6O2 = 6CO2 + 6H2O + ENERGY

    - CO2 production is equal to 200 mL/min (VCO2)

    - Increased metabolism is seen with:

      - Hyperthermia

      - Exercise

      - Hyperthyroidism

      - Shivering or seizures

      - High carbohydrate loading

      - Sepsis

    - Decreased metabolism is seen with:

      - Hypothermia

    - If CO2 production increases, alveolar ventilation will increase to maintain a normal PaCO2. If a person cannot respond by increasing their alveolar ventilation, CO2 will increase (PaCO2 and PETCO2).

      - Patients with limited ventilatory reserve.

      - Patients on a control mode of ventilation.

      - Patients seizing.

      - Patients who are sedated.

  - Circulation

    - PETCO2 is an excellent index of overall ventilation only if alveoli are being perfused.

    - A decreased perfusion in relation to ventilation will increase VQ ratios in the lung and is reflected as an increase in alveolar deadspace.  A widening of the P(a-ET)CO2 gradient will occur.

    - Increased Alveolar Deadspace may be seen in the following conditions:

      - Pulmonary embolism

      - COPD

      - Shock (hypovolemia or cardiogenic)

      - Over-distended alveoli (PEEP or CPAP)

  - Ventilation

    - CO2 diffuses so rapidly (20 times greater than O2) that the PaCO2 usually provides an excellent index of the adequacy of overall ventilation.  As blood perfuses the alveoli, the CO2 of the blood equilibrates with the alveolus and we are thus able to monitor changes in ventilation through the use of exhaled CO2 levels.

    - PETCO2 is a reflection of the ventilation of PERFUSED alveoli.

    - Alveolar Ventilation

      - Increase VA, decrease PaCO2, decrease PETCO2

      - Decrease VA, increase PaCO2, increase PETCO2

    - NOTE:  There are no real VQ abnormalities in pure hyper- or hypoventilation.  If the PaCO2 increases, PETCO2 will increase unless it is so severe that the patient is not breathing at all (respiratory arrest) or is only moving deadspace gas (ineffective ventilation). In this instance, the PETCO2 will sharply decrease.

- VQ RATIOS

  - Normal VQ Ratio

    - Alveolar Ventilation (VA) is 4.0 L.

    - Cardiac Output (Qt) is 5.0 L.

    - VQ ratio is 4/5 = 0.8.

  - Deadspace

    - Ventilation in excess of perfusion.

      - Excess of ventilation or decrease in perfusion.

    - Increases in alveolar deadspace will produce a High VQ ratio and this will increase the P(a-ET)CO2.

      - Areas in the lung with a high VQ ratio will have low exhaled CO2 levels that mix with areas of normal VQ ratios.  The result is a lower PETCO2 level and a widening of the P(a-ETt)CO2 gradient.

      - Conditions that will cause a widened gradient include:

        - COPD

        - Pulmonary hypoperfusion

        - Pulmonary embolism

        - Air embolism

        - Cardiac arrest

        - Shock (hypovolemia or cardiogenic)

    - Many studies have shown a good correlation between the Vd/Vt ratio and the P(a-ET)CO2.

  - Shunts

    - Perfusion in excess of ventilation.

      - No ventilation or increased perfusion.

    - A true shunt produces a VQ ratio of 0.

    - A relative shunt (shunt effect) produces a low VQ ratio.

    - Shunts or shunt effects will not usually affect the P(a-ET)CO2 gradient (gradient will be normal).

    - Shunts or shunt effects are caused from:

      - Atelectasis

      - Pneumonia

      - Mucus plugging

      - Bronchial intubation

- P(a-ET)CO2 Gradient

  - The normal gradient is less than 5 mm Hg.

  - To calculate a gradient, you must obtain an ABG and measure the PETCO2 when the blood gas is obtained.

  - The PETCO2 obtained from the capnograph is used to approximate the alveolar CO2.

    - EXAMPLE:

      - PaCO2 of 40 mm Hg.  PETCO2 is 30 mm Hg.

      - P(a-ET)CO2 gradient = 40 – 30 = 10 mm Hg.

  - When comparing PaCO2 with PETCO2 it is important to adjust the PaCO2 to the patient’s body temperature if it is not 37 C.

  - Failure to adjust the PaCO2 for temperature will result in an incorrect P(a-ET)CO2.

- USE OF PETCO2 MONITORING DURING CPR

  - PETCO2 will decrease to 0 during a cardiac arrest (PaCO2 will elevate).

  - With CPR, PETCO2 should increase.

  - The PETCO2 correlates with cardiac output (pulmonary blood flow) during CPR.

  - PETCO2 can be used as a prognosticator of successful resuscitation.

  - Higher PETCO2 levels were observed during resuscitation of animals that were successfully resuscitated than during the resuscitation of those that were not revived.

  - A low PETCO2 level during CPR, is not likely to be associated with a positive outcome.

- CLINICAL APPLICATIONS OF CAPNOGRAPHY - TROUBLESHOOTING

  - Changes in the PETCO2 readings may occur due to changes in patients condition or equipment malfunction (ET tube, ventilator or capnograph).

  - Sudden loss of PETCO2 to zero or near zero indicates immediate danger because no respiration is detected.

    - Esophageal intubation

    - Complete airway disconnection from ventilator

    - Ventilator malfunction/capnograph malfunction

    - Totally obstructed/kinked endotracheal tube

    - Patient apneic

      - CPAP or spontaneously breathing

    - No perfusion (cardiac arrest)

  - Exponential decrease in PETCO2

    - Decreasing Qt.

    - Sudden hypotension (hypovolemia, cardiogenic shock, hemorrhage)

    - Circulatory arrest with continued pulmonary ventilation.

    - Pulmonary embolism

  - Sustained low PETCO2

    - Hyperventilation

Example: metabolic acidosis, hypoxia, anxiety or pain.

    - Hypothermia

      - Decreased production of CO2.

    - Alveolar deadspace

      - Widened P(a-ET)CO2 gradient and loss of alveolar plateau.

        - COPD, partial airway obstruction from kinked ET tube or mucus plugging, bronchospasm.

      - Widened P(a-ET)CO2 gradient and good alveolar plateau.

        - Pulmonary embolism, hypovolemia, pulmonary hypoperfusion, Endotracheal tube in the hypopharynx or leak in the ET tube cuff (loss of waveform)

  - Increase in PETCO2

    - Increased metabolism (hyperthermia, seizures)

    - Hypoventilation (respiratory depressants, metabolic alkalosis, neuromuscular disease, muscular paralysis)

    - Leak in ventilator circuit (patient is hypoventilated)

  - Cleft in the alveolar plateau

    - A downward spike ("curare cleft") may be seen the patient recovering from neuromuscular blockade.

    - The patient may not be breathing in synchrony with the ventilator.

    - Should not be exclusively associated with neuromuscular blockage

  - A rise in the baseline

    - Patient is rebreathing CO2 (malfunction of exhalation valve, CO2 absorber)

    - Capnograph needs recalibrating

  - Cardiac Oscillations

    - In patients with bradypnea, cardiac oscillations may be seen on the capnogram and is the result of the heart beating against the lungs.

- Calibration:

  - High Point Calibration 5% CO2

  - Low Point Calibration can use room air which contains 0.3 mm Hg or .04% CO2.

- This is assumed to be zero.